Product 1430295 V1000 keypad assembly	Y/N
Personnel involved / Steve Nixon, Derek Lamb, Jonathan Connor, Michael Green, Phil Crossley / Danielson Schurter personnel: Darren Rhodes, Arron Russell
Reason for Change / Potential for moisture being trapped within the part, as the small area of exposed silver along the base of the tail of affected units show signs of corrosion on the powered track, which eventually causes the parts to fail / The problem looks to be that moisture is attacking the exposed track, I believe that this is exacerbated by there being no adhesive along the lower edge of the membrane. This is probably where moisture (ultrasound gel/cleaning fluid) is getting in. The keypad design has passed functional tests and the tracks are adequate for the power consumption.
Proposed design change & action to be taken / Design change not implemented yet. In the meantime, Michael Green is preparing revised service and production procedures.
Change of functional & action to be taken / No change to function.
Change to performance & action to be taken / Reliability improved by applying a small bead of silicone to the exposed areas of the track. This applies to units being service or produced. This is intermediate action.
Change of usability & action to be taken / No change to usability.
Change in safety features & action to be taken / No change to safety features.
Application of new standards / None
Does this have an effect on
Constituent parts	N
Product in process	N
Product already delivered	N
Inputs of risk management	N
Outputs of risk management	N
Product realization	N
Risk Analysis on any design change Refer QC22	N
Signed S. Nixon 22nd November 2017