NOTE: CapnoTrue®AMP/ASP/MG monitors do not require a routine Safety Check - STK (§6) or routine Measurement Accuracy Check - MTK (§11) according to German regulation (Attachment 1 and 2 of MPBetreibV)

The Handheld capnograph and pulse oximeter CapnoTrue®AMP/ASP/MG with accessories is permanently factory calibrated. The maintenance and calibration-free technology (gas analyzer and SpO2 Module) integrated in the CapnoTrue®AMP/ASP/MG ensures a robust measurement function throughout the lifetime of the monitor. No routine calibration is required however a basic maintenance plan is highly recommended.

To document the maintenance, use the following checklist. All recommended test steps are described in detail in the document CapnoTrue Maintenance instructions and Service Manual.

| Description of device components |  |  |  |  |
| --- | --- | --- | --- | --- |
| Device | Device Model (REF on back of label) CapnoTrue®AMP (3090112001) CapnoTrue®MG (3090112009) CapnoTrue®ASP (3090112002) |  |  |  |
| External power supply | black connector at cable and at device (Rev. 1) orange connector at cable , metal connector at device (Rev. 2) |  |  |  |
| SpO2-Sensor | SC 6500 | SCP 6500 | W 6500 |  |
|  | SCA 6500 | SCPA 6500 | Disposable |  |
| Serial number (SN) Of the modules in the device menu: Service Menu> System information > Device information | Device (SN on rear device label): |  |  |  |
|  | SpO2-Sensor (SN on cable): |  |  |  |
|  | SpO2 Module (SN in Service Menu): |  |  |  |
|  | CO2 Module (SN in Service Menu or on housing): |  |  |  |
| Firmware Version update at manufacturer is possible starting from v7.5 | Version (Startup-Screen at switch on): |  |  |  |

| Comments |  |  |
| --- | --- | --- |
| Result | Pass AMP/MG: Step 1 – 5 ASP: Step 1 - 6 | Fail Don’t use device to monitor patients! Perform repair. |
| Date | Tester |  |

| Test step | Target state | Result |  | comment |
| --- | --- | --- | --- | --- |
| 1. Visual inspection |  |  |  |  |
| Housing, Display, Keypad | No damages or defects | pass | fail |  |
| Sensors and cable | No exposed optical components or cables or other damages to the sensor | pass | fail |  |
| external power supply (cable+ connector) | No damages or defects | pass | fail |  |
| Label + transparent protective foil | Label and the transparent protective foil are fixed clean to the back of the device. | pass | fail |  |
| Screw lock | Fixed to upper right-hand screw | pass | fail |  |
| 2. Power Supply Options (AA batteries, Li-Poly rechargeable battery, external power supply) (switch on with each supply option >Measurements: SpO2 at finger; etCO2 by breathing through airway adapter > switch off) |  |  |  |  |
| only 4 AA Batteries | switch on> measurement> switch off | pass | fail |  |
| only Li-Poly | switch on> measurement> switch off | pass | fail |  |
| only ext. power supply | switch on> measurement> switch off | pass | fail |  |
| ext. power supply + Li-Poly | switch on > measurement + charging > switch off | pass | fail |  |
| Comments: (Pixel defects, etc.) |  |  |  |  |
| 3. SpO2-Function at simulator |  |  |  |  |
| SpO2 at simulator setpoint 82% (BPM or Nellcor curve) | 82% SpO2 (+-2 digits) | pass | fail | value: |
| Pulse rate at simulator setpoint 75bpm | 75 beats/min (+-1 digits) | pass | fail | value: |
| SpO2 low alarm | Audible and visual | pass | fail |  |
| Comments: |  |  |  |  |
| 4. CO2-Measurement (use reference device as alternative means if no breath simulator is available) Select Mode at CapnoTrue: Main Menu > Service > Maintenance > Check gas accuracy Calibration gas (5% ±0.1vol% CO2, 20.9% O2 balance N2) |  |  |  |  |
| 5vol% CO2 calibration gas + airway adaptor | EtCO2:4.7 – 5.3 vol% FiCO2: 0.0 – 0.3vol% | pass | fail | value: EtCO2 / FiCO2 |
| Breath per minute | setpoint (+-1 digit) | pass | fail | value: |
| Comments:: |  |  |  |  |
| 5. Alarms und Data Download (switch on device, SpO2 Sensor at finger, breath through airway adapter) |  |  |  |  |
| Remove SpO2 Sensor from finger | Alarm sound and message „SpO2 Probe off !“ | pass | fail |  |
| Breath 3 X through airway adapter and wait 20s | Alarm sound and message „Apnoea !!“ | pass | fail |  |
| Download data to PC with CapnoTrue PC-Software | Successful data download to PC | pass | fail |  |
| Delete all data | All stored data in the device are deleted | pass | fail |  |
| Comments: |  |  |  |  |
| 6. Check pressure – ISA module (Only for sidestream capnograph CapnoTrue ASP) Select at device: Main Menu > Service > Maintenance > Check pressure |  |  |  |  |
| Compare displayed ambient pressure value with the actual barometric pressure value at the local site | actual barometric pressure ±5 kPa | pass | fail |  |
| Compare displayed cuvette pressure with ambient pressure | cuvette pressure is 1 to 5 kPa smaller than displayed ambient pressure | pass | fail | Ambient pr. kPa Cuvette pr. kPa |
| Comments: |  |  |  |  |