NOTE: CapnoTrue®AMP/ASP/MG monitors do not require a routine Safety Check - STK (§6) or routine Measurement Accuracy Check - MTK (§11) according to German regulation (Attachment 1 and 2 of MPBetreibV)

The Handheld capnograph and pulse oximeter CapnoTrue®AMP/ASP/MG with accessories is permanently factory calibrated. The maintenance and calibration-free technology (gas analyzer and SpO2 Module) integrated in the CapnoTrue®AMP/ASP/MG ensures a robust measurement function throughout the lifetime of the monitor. No routine calibration is required however a basic maintenance plan is highly recommended.

The following maintenance tasks shall be performed:

1)	Gas readings should be verified at regular intervals with a reference instrument or with calibration gas (5% ±0.1vol% CO2, 20.9% O2 balance N2). The suggested interval for gas span check is once every year and whenever gas readings are questionable.

2)	Perform a functional SpO2 test using a SpO2 simulator. However note that a functional tester cannot be used to assess the accuracy of a pulse oximeter probe or a pulse oximeter monitor, only its functionality. The only reliable method of testing the SpO2 measurement accuracy of the monitor is the clinical validation of the measurement data, indicated by the system monitor with SpO2 sensor on the basis of a blood gas analysis. During extensive clinical studies, the monitor combined with the approved sensors evidenced the accuracy required. Verify SpO2 readings at regular intervals with a reference instrument.

3)	Test of the alarm system. In order to trigger an alarm for test purposes during monitoring set the upper alarm limit of SpO2 or pulse rate below the currently indicated measurement value. The device will react with a visual and audible alarm.

4)	Check the data download function at a PC with the PC-software

5)	Perform a visual inspection of all cables accessories and housing parts

6)	Perform a functional check of all power supply options (Li-Poly battery, Mignon AA battery, external power supply)

To document the maintenance, use the following checklist. All recommended test steps are described in detail in the document CapnoTrue Maintenance instructions and Service Manual.

| Description of device components for documentation of the maintenance of the device and its components throughout the live time of the device. |  |  |  |  |
| --- | --- | --- | --- | --- |
| Device Identify device | Device Model (REF on back of label) CapnoTrue®AMP (3090112001) CapnoTrue®MG (3090112009) CapnoTrue®ASP (3090112002) |  |  |  |
| External power supply documentation of revision | black connector at cable and at device (Rev. 1) orange connector at cable , metal connector at device (Rev. 2) |  |  |  |
| SpO2-Sensor Sensor type is noted on the label fixed to the sensor cable | SC 6500 SCP 6500 |  | W 6500 |  |
|  | SCA 6500 | SCPA 6500 | Disposable |  |
| Serial number Of the SpO2 module in the device menu: Service Menu> System information > Device information Changing the default start up settings of the device can seriously affect the functionality of the device and the alarms! Only suitably qualified personnel with the required technical and medical knowledge are allowed to change the default start up settings! | Device (SN on rear device label): |  |  |  |
|  | SpO2-Sensor (SN on cable): |  |  |  |
|  | SpO2 Module (SN in Service Menu): |  |  |  |
|  | CO2 Modul (SN in Service Menu or on housing): |  |  |  |
| Firmware Version startin from v7.5 a software upgrade is available at bluepoint Medical | Version (Startup-Screen at switch on): |  |  |  |

| Comment | Detailed description on repairs performed to pass the functional test etc. |  |
| --- | --- | --- |
| Result | Pass AMP/MG: Test 1 – 5 ASP: Test 1 – 6 Tick if all tests successfully passed | Fail Don’t use device to monitor patients! Perform repair. Send device to SERVICE of bluepoint MEDICAL or trained and certified service partners |
| Date | Tester |  |

| Test step | Target state | Result |  | comment |
| --- | --- | --- | --- | --- |
| 1. Visual inspection |  |  |  |  |
| Housing, Display, Keypad | No damages or defects | pass | fail |  |
| Sensors and cable | No exposed optical components or cables or other damages to the sensor | pass | fail |  |
| external power supply (cable+ connector) | No damages or defects | pass | fail |  |
| Label + transparent protective foil | Label and the transparent protective foil are fixed clean to the back of the device. | pass | fail |  |
| Screw lock | Fixed to upper right-hand screw | pass | fail |  |
| 2. Power Supply Options (AA batteries, Li-Poly rechargeable battery, external power supply) (switch on with each supply option >Measurements: SpO2 at finger; etCO2 by breathing through airway adapter > switch off) |  |  |  |  |
| only 4 AA Batteries | switch on> measurement> switch off | pass | fail |  |
| only Li-Poly | switch on> measurement> switch off | pass | fail |  |
| only ext. power supply | switch on> measurement> switch off | pass | fail |  |
| ext. power supply + Li-Poly | switch on > measurement + charging > switch off | pass | fail |  |
| Comments: (Pixel defects, etc.) |  |  |  |  |
| 4 x AA Batteries, or 1 x Li-Poly:external power supply: Before measurement ensure that AA batteries are full and Li-Poly charged |  |  |  |  |
| 3. SpO2-Function at simulator |  |  |  |  |
| SpO2 at simulator setpoint 82% (BPM or Nellcor curve) e.g. INDEX II simulator (Biotek, Fluke) select the SpO2 cal. Curve BPM or NELLCOR under MAKE | 82% SpO2 (+-2 digits) NOTE: SpO2 low alarm sounds, SpO2 measurement value turns yellow and is flashing. | pass | fail | value: write down displayed SpO2 value |
| Pulse rate at simulator setpoint 75bpm set pulse rate at simulator | 75 beats/min (+-1 digits) | pass | fail | value: write down displayed pulse rate value |
| SpO2 low alarm | Audible and visual | pass | fail |  |
| Comments: |  |  |  |  |
| 4. CO2-Measurement (use reference device as alternative means if no breath simulator is available) Select Mode at CapnoTrue: Main Menu > Service > Maintenance > Check gas accuracy Calibration gas (5% ±0.1vol% CO2, 20.9% O2 balance N2) Verify gas accuracy reading in Test Screen via SERVICE MENU: - Switch on device - connect new IRMA airway adapter (or sampling line for ASP), perform zeroing - Select at device: Main Menu > Service > Maintenance > Check gas accuracy - Apply calibration gas to airway adapter and verify reading |  |  |  |  |
| 5vol% CO2 calibration gas + airway adaptor e.g. | CO2 should be Gas conc. (+- (0.2 vol%. + 2% of measurement value) At 5vol% CO2 the limits are: EtCO2:4.7 – 5.3 vol% FiCO2: 0.0 – 0.3vol% | pass | fail | value: EtCO2 / FiCO2 write down displayed CO2 values |
| Breath per minute connect airway adaptor to adaptor of reference device. Breath uniformly through adaptors for 30 minutes. Compare values | reference value (+-1 digit) + tolerance of reference device | pass | fail | value: write down displayed respiration rate at CapnoTrue and reference device |
| Comments:: |  |  |  |  |
| 5. Alarms und Data Download (switch on device, SpO2 Sensor at finger, breath through airway adapter) |  |  |  |  |
| Remove SpO2 Sensor from finger | Alarm sound and message „SpO2 Probe off !“ 2 x beep every 16s | pass | fail |  |
| Breath 3 X through airway adapter and wait 20s | Alarm sound and message „Apnoea !!“ 3 x beep every 5s | pass | fail |  |
| Download data to PC with CapnoTrue PC-Software 1. Connect device via USB to PC 2. Verify message “USB connected at device 3. Open CapnoTrue SW 4. download data according to instructions at “communication download” | Successful data download to PC Compare data in Tab Diagram and stored alarms in Tab Patient info with data at device - If errors occur during download check compatibility of versions (see page 3 of Instructions for Use) - check if correct COM port is selected - check if the USB drivers were installed according to the instructions | pass | fail |  |
| Delete all data Select Delete all data under: Menu > data management > delete all data | All stored data in the device are deleted | pass | fail |  |
| Comments: |  |  |  |  |
| 6. Check pressure – ISA module (Only for sidestream capnograph CapnoTrue ASP) Select at device: Main Menu > Service > Maintenance > Check pressure |  |  |  |  |
| Compare displayed ambient pressure value with the actual barometric pressure value at the local site | actual barometric pressure ±5 kPa | pass | fail |  |
| Compare displayed cuvette pressure with ambient pressure | cuvette pressure is 1 to 5 kPa smaller than displayed ambient pressure | pass | fail | Ambient pr. kPa Cuvette pr. kPa write down displayed values |
| Comments: |  |  |  |  |