
Use of VM2160 with SPO2 Simulator Index 2

Simulator: Index 2      Manufacturer: Fluke / Bio-Tek

Software 3.00 

Curve/Make: BCI

AMP 3.000 %

TLC 1100
Ambient Normal

	Index 2
BCI Curve
Pulse 75
Saturation  [ % ]


	Reading on VM2160
with SC6500
[ % ]

	 
	min
	typical
	max

	100
	97,02
	99,00
	99,00

	99
	97,02
	99,00
	99,00

	98
	96,04
	98,00
	99,00

	97
	95,06
	97,00
	98,94

	96
	94,08
	96,00
	97,92

	95
	93,10
	95,00
	96,90

	94
	92,12
	94,00
	95,88

	93
	91,14
	93,00
	94,86

	92
	89,18
	91,00
	92,82

	91
	88,20
	90,00
	91,80

	90
	87,22
	89,00
	90,78

	89
	86,24
	88,00
	89,76

	88
	85,26
	87,00
	88,74

	87
	84,28
	86,00
	87,72

	86
	83,30
	85,00
	86,70

	85
	82,32
	84,00
	85,68

	84
	81,34
	83,00
	84,66

	83
	80,36
	82,00
	83,64

	82
	79,38
	81,00
	82,62

	81
	78,40
	80,00
	81,60

	80
	77,42
	79,00
	80,58

	79
	76,44
	78,00
	79,56

	78
	75,46
	77,00
	78,54

	77
	74,48
	76,00
	77,52

	76
	73,50
	75,00
	76,50

	75
	72,52
	74,00
	75,48

	74
	71,54
	73,00
	74,46

	73
	70,56
	72,00
	73,44

	72
	69,58
	71,00
	72,42

	71
	68,60
	70,00
	71,40

	70
	67,62
	69,00
	70,38


[image: image1.jpg]


Please note that according to the international standard ISO 9919:2005, chapter 50.101.2 the use of patient simulators or functional testers shall not be used to validate the accuracy of pulse oximeter equipment!

The only reliable method of testing the measurement accuracy of an SpO2 monitor  is  the clinical validation of the measurement data, indicated by the system monitor + sensor on the basis of  blood gas analysis. 

During extensive clinical studies, the VM2160  combined with the approved sensors evidenced the accuracy required.

06/10/09[image: image2.png]




