	     [image: image1.png]



	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley -West Yorkshire BD20 7DT - United Kingdom
Tel: +44 1535 634542                 Fax: +44 1535 635582

Email: info@viamed.co.uk  Website: www.viamed.co.uk



REV.0-V
Maintenance Checklist
VM-2500 Series Capnographs

The Viamed VM-2500 Series capnographs and pulse oximeters are permanently factory calibrated. The maintenance and calibration-free technology (gas analyzer and SpO2 module) integrated into the VM-2500 Series ensures a robust measurement function throughout the lifetime of the monitor and no routine calibration is required. However, a basic maintenance plan is highly recommended. 

To document the maintenance, use the following checklist. 
	Details of the device components

	Device
	Device Model (REF on rear label)
( VM-2500-S   ( VM-2500-M   ( VM-2500-MG

	External power supply
	( black connector on the cable and on the device (Rev. 1) 

( orange connector on the cable, metal connector on the device (Rev. 2)

	SpO2 Sensor


	( SC 6500
	( SCP 6500
	( W 6500
	( Disposable

	Serial Number (SN)
The serial numbers of the modules can be found in the device menu: Service Menu> System information > Device information


	Device (SN on device rear label):

	
	SpO2-Sensor (SN on cable):

	
	SpO2 Module (SN in Service Menu):

	
	CO2 Module (SN in Service Menu or on device housing):

	Firmware Version

Factory updates are possible, starting from v7.5
	Version (Start-up-Screen at switch on):


	Comments
	

	Result
	 Pass
VM-2500-M & VM-2500-MG: Steps 1 – 5

VM-2500-S: Steps 1 - 6
	 Fail
Don’t use the device to monitor patients! Perform repair.

	Date


	Tester


	Test step
	Target state
	Result
	comment

	1. Visual inspection

	Housing, Display, Keypad
	No damages or defects
	 pass
	 fail
	

	Sensors and cable
	No exposed optical components, wires or other damage to the sensor
	 pass
	 fail
	

	External power supply (cable & connector)
	No damage or defects
	 pass
	 fail
	

	Label & transparent protective foil
	Label and the transparent protective foil are intact. Clean and affixed to the rear of the device.
	 pass
	 fail
	

	Screw lock
	Fixed to upper right-hand screw
	 pass
	 fail
	

	2. Power Supply Options (AA batteries, Li-Poly rechargeable battery, external power supply)
(switch on with each supply option - confirm measurement readings: SpO2 at finger; etCO2 by breathing through airway adapter -  switch off)

	Only 4 AA Batteries
	switch on – measurement - switch off
	 pass
	 fail
	

	Only Li-Poly
	switch on – measurement - switch off
	 pass
	 fail
	

	Only ext. power supply
	switch on – measurement - switch off
	 pass
	 fail
	

	External power supply 
& Li-Poly
	switch on – measurement - switch off
	 pass
	 fail
	

	Comments: 
(Pixel defects, etc.)
	

	3. SpO2 - Function testing with a simulator

	SpO2 at simulator setpoint 82%
(Nellcor R’ curve)
	82% SpO2 (+-2 digits)
	 pass
	 fail
	value:

	Pulse rate at simulator setpoint 75 bpm
	75 beats/min (+-1 digits)
	 pass
	 fail
	value:

	SpO2 low alarm
	Audible and visual
	 pass
	 fail
	

	Comments:
	

	4. CO2 - Measurement (If a breath simulator is not available then alternatively use a reference device)

Select Mode on the VM-2500: Main Menu > Service > Maintenance > Check gas accuracy
Calibration gas (5% ±0.1vol% CO2, 20.9% O2 balance N2)

	5vol% CO2 calibration gas + airway adaptor
	EtCO2:4.7 – 5.3 vol%

FiCO2: 0.0 – 0.3vol%
	 pass
	 fail
	value: EtCO2 / FiCO2

	Breaths per minute
	setpoint (+-1 digit)
	 pass
	 fail
	value:

	Comments:

	

	5. Alarms & Data Download (switch on device with SpO2 Sensor attached to a finger, breath through airway adapter)

	Remove SpO2 Sensor from finger
	Alarm sound and message
“SpO2 Probe off !“
	 pass
	 fail
	

	Breath 3 times through airway adapter and wait 20s
	Alarm sound and message
“Apnoea !!“
	 pass
	 fail
	

	Download data to PC with VM-2500 PC-Software
	Successful data download to PC
	 pass
	 fail
	

	Delete all data
	All stored data in the device is deleted
	 pass
	 fail
	

	Comments:
	
	
	
	

	6. Check pressure – ISA module (Only for sidestream capnograph VM-2500-S)

On the device select: Main Menu > Service > Maintenance > Check pressure

	Compare displayed ambient pressure value with the actual local barometric pressure value 
	actual barometric pressure
 ±5 kPa 
	 pass
	 fail
	

	Compare displayed cuvette pressure with ambient pressure
	cuvette pressure is 1 to 5 kPa smaller than displayed ambient pressure
	 pass
	 fail
	Ambient pressure:    
                          kPa           

Cuvette pressure:

                          kPa

	Comments:
	
	
	
	


[image: image1.png]