Post Market Surveillance Review

4310001 Flowsensor SpiroTrue A. Box of 5

First sold: 12th January 2009

General sales suspended 18th December 2013 due to MHRA stance on single use designation.

Total sales to date: 5,906 boxes, across 34 countries

Total failures: 6 boxes

Total returns: 6 boxes

Total customer failures: 0.102%

Liverpool Heart & Chest Hospital SRS63066, received 1 box 10th October 2011

Failed calibration. Also reported sensors generally have shorter life than Draeger OEM sensors.

Result: Customer lost, approximately 100 boxes per year. Resumed using Draeger OEM sensors, will try and regain the sales with Flowsensor SpiroTrue A.

Dorset County Hospital SRS63869, received 1 box 3rd June 2013

Failed calibration.

Result: Customer lost, approximately 12 boxes per year. Will try and regain the sales with Flowsensor SpiroTrue A.

Papworth Hospital SRS64002, received 4 boxes 11th September 2013

Failed calibration when used on Evita 4s.

Result: Customer lost, approximately 120 boxes per year. They have tried both SpiroTrue A and SpiroTrue A (PC). Will try and regain the sales with Flowsensor SpiroTrue A.

Applications:

| Manufacturer | Description |  |
| --- | --- | --- |
| Acutronic | 1601+ | Flow Sensor |
| Acutronic | SpiroQuant A | Flow Sensor |
| Asmuth GmbH Medinzintechnik | AS432+ | Flow Sensor |
| Asmuth GmbH Medinzintechnik | AS8403735 | Flow Sensor |
| Asmuth GmbH Medinzintechnik | SpiroTrue A | Flow Sensor |
| Draeger | Spirolog 84 03 735 | Flow Sensor |
| Draeger | 8403735 | Flow Sensor |
| Draeger | Anemone | Monitor |
| Draeger | Apollo | Anaesthesia System |
| Draeger | AV1 | Anaesthesia System |
| Draeger | Cato | Anaesthesia System |
| Draeger | Cicero | Anaesthesia System |
| Draeger | D-9118-2009 | Flow Sensor |
| Draeger | Dura | Ventilator |
| Draeger | Eva | Ventilator |
| Draeger | Evita | Ventilator |
| Draeger | Evita 4 | Ventilator |
| Draeger | Evita XL | Ventilator |
| Draeger | Fabius GS | Anaesthesia System |
| Draeger | Infinity ID (NOTE Viamed does not have RFID) | Flow Sensor |
| Draeger | Infinity V500 (NOTE Viamed does not have RFID) | Anaesthesia System |
| Draeger | Julian | Anaesthesia System |
| Draeger | MT-2086-2007 (NOTE Viamed does not have RFID) | Flow Sensor |
| Draeger | PM 8030 | Monitor |
| Draeger | PM 8035 | Monitor |
| Draeger | PM 8050 | Monitor |
| Draeger | PM 8060 | Monitor |
| Draeger | Primus | Anaesthesia System |
| Draeger | Savina | Ventilator |
| Draeger | Spirolog | Flow Sensor |
| Draeger | Spirolog 8403735 | Flow Sensor |
| Draeger | Sulla | Anaesthesia System |
| Envitec | 1000470 | Flow Sensor |
| Envitec | SpiroQuant A+ | Flow Sensor |
| Integral Process | 0032 00 02 | Flow Sensor |
| Integral Process | 65025 | Flow Sensor |
| Integral Process | 8403735 | Flow Sensor |
| Intersurgical | 2777000 | Flow Sensor |
| Intersurgical | e-FlowSensor | Flow Sensor |
| Medacx | M-45-00-0005 | Flow Sensor |
| Medacx | SpiroQuant A | Flow Sensor |
| Proact Medical | 1000470 | Flow Sensor |
| Proact Medical | SpiroQuant A+ | Flow Sensor |

Reviewed by: Steve Nixon

Director – Viamed Ltd.

Date: 21st July 2017