S Hardaker – 6th September 2007

# Apgar Timer Software Revision v2.01

## Objective

At the request of Peter Anderson, I evaluated the latest Apgar Timer with software revision v2.01 with regards to functionality.

## Method

The following tests were performed:

- Using a matrix of all possible key-press combinations in each mode, verified that all of the keys performed the expected action in all situations.

- Verified that all expected actions occur at the expected times, including the timing intervals and format of the audible indications.

- Verified that the timer runs for a duration of 1 hour, reaching 59:59 on the display and then reverting back to stand-by mode. This test performed twice.

- Verified that the timer keeps time with other timing devices, I used a stop-clock and wristwatch (note: neither devices certified as calibrated, but all 3 timing devices remained synchronised as far as is detectable by the eye).

## Results

| MODE | ### KEY-PRESS OR ACTION | EXPECTED ACTION | ACTUAL ACTION | OK? (Y/N) |
| --- | --- | --- | --- | --- |
| Key-presses |  |  |  |  |
| Stand-by Mode | Key – Start | Enters counting mode | Enters counting mode | Y |
| Stand-by Mode | Key - Freeze/Unfreeze | No action | No action | Y |
| Stand-by Mode | Key – Reset | No action | No action | Y |
| Counting Mode | Key – Start | No action | No action | Y |
| Counting Mode | Key - Freeze/Unfreeze | Freezes count | Freezes count | Y |
| Counting Mode | Key – Reset | Enters stand-by mode | Enters stand-by mode | Y |
| Freeze Mode | Key – Start | Continues real-time count | Continues real-time count | Y |
| Freeze Mode | Key - Freeze/Unfreeze | Continues real-time count | Continues real-time count | Y |
| Freeze Mode | Key – Reset | Enters stand-by mode | Enters stand-by mode | Y |
| Actions |  |  |  |  |
| Off (no batteries inserted) | Insert batteries | Version 2.01 displayed briefly, then enters stand-by mode. | Version 2.01 displayed briefly, then enters stand-by mode. | Y |
| Counting Mode | Action – Display reaches 1:00 | 1 audible beep | 1 audible beep | Y |
| Counting Mode | Action – Display reaches 5:00 | 2 audible beeps | 2 audible beeps | Y |
| Counting Mode | Action – Display reaches 10:00 | 3 audible beeps | 3 audible beeps | Y |
| Counting Mode | Action – Display reaches 59:59 | Enters stand-by mode | Enters stand-by mode | Y |

## Conclusion

As far as I am able to determine, the software performs as per the design specification in terms of functionality.

I have not been able to verify that the software modifications to extend battery life have achieved their intended aim, only that there has been no obvious effect on the functionality.

SJH – 6/9/07