## Sheet1
Product | Certificate Reference | Internal Issue ID | Non Conformance ID | Non Conformances Broken out | Route Cause | Correction Action Plan | CAP TimeScale | CAP Responsibility
Oxygen Sensors | CE 97289 | 101076 | 1515871-201705-M1 | “Application of the classification rules shall be governed by the intended purpose of the devices. If the device is intended…”
Oxygen Sensors | CE 97289 |  |  | - No mention of device classification was provided in the technical documentation, or in the essential requirements. No Declaration of conformity declaring such compliance was provided. | 1. Fault with PDF version of the Essential Requirement 2. File Exporter only Exporting Files directly tagged to Specific Technical files – The generic files that are used in multiple Technical files were not exporting. Reason not detected before. First remote audit. And First time PDF of Essential requirements used – digital on screen view lists all files... | 1. PDF output to be redesigned to list all documents in relevant ER sections 2. recode the exporter so it also exports the more generic documents attached to multiple Technical Files. 3 Explanatory documents and notes for this section will be added if applicable | 2017-11-01 | Derek Lamb
Oxygen Sensors | CE 97289 |  |  | It is unclear that the device is properly classified (IIa vs IIb) based on intended use statements.
Oxygen Sensors | CE 97289 | 101080 | 1515871-201705-M2 | The devices must be designed and manufactured in such a way that, when used under the conditions and for the purposes intended, they will not compromise the clinical condition or the safety of patients, or the safety and health of users or, where applicable, other persons…
Oxygen Sensors | CE 97289 |  |  | [ER 4] - While the risk management files clear states there is a finite functional lifetime for the device, “Oxygen Sensors measure oxygen, as they do so they use up the electrolyte,” no shelf life testing or functional life testing has been provided. No evaluation or derivation of such a lifetime has been provided. | 1. Fault with PDF version of the Essential Requirement 2. File Exporter only Exporting Files directly tagged to Specific Technical files – The generic files that are used in multiple Technical files were not exporting. Reason not detected before. First remote audit. And First time PDF of Essential requirements used – digital on screen view lists all files... | 1. PDF output to be redesigned to list all documents in relevant ER sections 2. recode the exporter so it also exports the more generic documents attached to multiple Technical Files. 3 Explanatory documents and notes for this section will be added if applicable | 2017-11-01 | Derek Lamb
Oxygen Sensors | CE 97289 |  |  | [ER 5, 8.6] - No evidence of packaging validation or evidence that the device is shipped within the design limits specified by the OEM gave been provided. No visual inspection or device function testing has been provided post Shipment.
Oxygen Sensors | CE 97289 |  |  | [ER 7.3] - The ERC states explicitly the device comes into contact with oxygen and other anesthetic gases. While evidence for accuracy is supplied by EN 80601-2-55 testing explicit evidence for use in combination with other foreseeable gases has not been provided. Risk associated with said contact has not been evaluated.
Oxygen Sensors | CE 97289 |  |  | [ER 9.1] - A large list of available devices into which the sensor is installed is provided in the instructions for use. However the technical documentation provided while supplying technical limits to the sensors does not demonstrate in the IFU or in design verification that the device will function when used as Indicated.
Oxygen Sensors | CE 97289
Oxygen Sensors | CE 97289 | 101084 | 1515871-201705-M3 | The solutions adopted to fulfill the essential requirements are incomplete. The system presented systemically fails to follow ISO 147971 requirements, including but not limited to:
Oxygen Sensors | CE 97289 |  |  | assignment of responsibilities and authorities per ISO 14971 section 3.4b
Oxygen Sensors | CE 97289 |  |  | review of risk management activities per ISO 14971 section 3.4c
Oxygen Sensors | CE 97289 |  |  | includes verification activities per ISO 14971 section 3.4e
Oxygen Sensors | CE 97289 |  |  | includes activities related to production and post-production information collection and review per ISO 14971 sections 3.4f and 9 While the risk management is recently updated there is no reference to post market surveillance data or evidence of maintenance of the risk management file in light of field data for the devices in question.
Oxygen Sensors | CE 97289 |  |  | Risks are not calculated both before and after risk controls (per 14971 section 6.4)
Oxygen Sensors | CE 97289 |  |  | Key relevant risks are identified as not applicable, with insufficient or incorrect rationales, including exclusion of monitoring functions as relevant risks.
Oxygen Sensors | CE 97289
Oxygen Sensors | CE 97289 | 101089 | 1515871-201705-M4 | The solution adopted to fulfil the Essential Requirements is not complete. The devices must achieve the performances intended by the manufacturer and be designed, manufactured and packaged in such a way that they are…
Oxygen Sensors | CE 97289 |  |  | Despite providing OEM function specifications, evidence of design verification tied to these design inputs was not provided beyond EN 80601-2- 55 safety testing. Furthermore design inputs by the manufacturer through which the components were selected IE. requirements which these components must meet, were not provided
Oxygen Sensors | CE 97289 |  |  | Testing against design inputs was not provided beyond standards Assessments. | 1. Fault with PDF version of the Essential Requirement 2. File Exporter only Exporting Files directly tagged to Specific Technical files – The generic files that are used in multiple Technical files were not exporting. Reason not detected before. First remote audit. And First time PDF of Essential requirements used – digital on screen view lists all files... | 1. PDF output to be redesigned to list all documents in relevant ER sections 2. recode the exporter so it also exports the more generic documents attached to multiple Technical Files. 3 Explanatory documents and notes for this section will be added if applicable | 2017-11-01 | Derek Lamb
Oxygen Sensors | CE 97289
Oxygen Sensors | CE 97289 | 101093 | 1515871-201705-M5 | The system for post-market clinical follow-up is inadequate. The clinical evaluation and its documentation must be actively updated with data obtained from the post-market surveillance. Where post-market clinical follow-up as part of the post-market surveillance plan for the device is not deemed necessary, this must be duly justified and documented Objective evidence
Oxygen Sensors | CE 97289 |  |  | While the supplied PMS document ID20910 was reviewed and found to contain data on field failures and resulting evaluation and action, no explicit PMCF or justification for not completing PMCF has been provided
Oxygen Sensors | CE 97289
Oxygen Sensors | CE 97289 | 101097 | 1515871-201705-M6 | The EC declaration of conformity does not fully meet the requirements of Annex II. The EC declaration of conformity is the procedure whereby the manufacturer who fulfils the obligations imposed by Section 1
Oxygen Sensors | CE 97289 |  |  | The manufacturer has failed to provide a Declaration of Conformity per Annex II and the NB-MED Consensus Statement S/01/99 | 1. Fault with PDF version of the Essential Requirement 2. File Exporter only Exporting Files directly tagged to Specific Technical files – The generic files that are used in multiple Technical files were not exporting. Reason not detected before. First remote audit. And First time PDF of Essential requirements used – digital on screen view lists all files... | 1. PDF output to be redesigned to list all documents in relevant ER sections 2. recode the exporter so it also exports the more generic documents attached to multiple Technical Files. 3 Explanatory documents and notes for this section will be added if applicable | 2017-11-01 | Derek Lamb
Oxygen Sensors | CE 97289
Oxygen Sensors | CE 97289 | 101100 | 1515871-201705-N1 | The solutions adopted to fulfill the essential requirements are incomplete. The devices must be designed and manufactured in such a way that, when used under the conditions and for the purposes intended, they will not compromise the clinical condition or the safety
Oxygen Sensors | CE 97289 |  |  | While the ERC ID21171 properly discusses that usability largely does not apply to the device in question as a component only, usability covers the full life of the product including installation. No evidence has been provided for who the end user would be (technicians) or device installation usability beyond basic review of risk of calibration.
Oxygen Sensors | CE 97289
Oxygen Sensors | CE 97289 | 101105 | 1515871-201705-N2 | The information supplied by the manufacturer does not fully meet the essential requirements
Oxygen Sensors | CE 97289 |  |  | Storage and shipment conditions between labels and IFU are not consistent. Specifically:
Oxygen Sensors | CE 97289 |  |  | Labeling clearly states handling temperature “5-20C” | 1. Fault with PDF version of the Essential Requirement 2. File Exporter only Exporting Files directly tagged to Specific Technical files – The generic files that are used in multiple Technical files were not exporting. Reason not detected before. First remote audit. And First time PDF of Essential requirements used – digital on screen view lists all files... | 1. PDF output to be redesigned to list all documents in relevant ER sections 2. recode the exporter so it also exports the more generic documents attached to multiple Technical Files. 3 Explanatory documents and notes for this section will be added if applicable | 2017-11-01 | Derek Lamb
Oxygen Sensors | CE 97289 |  |  | IFU states “Recommended storage 5-15C” and “Storage -20-50C” neither of which is in line with OEM requirements or labeling. | 1. Fault with PDF version of the Essential Requirement 2. File Exporter only Exporting Files directly tagged to Specific Technical files – The generic files that are used in multiple Technical files were not exporting. Reason not detected before. First remote audit. And First time PDF of Essential requirements used – digital on screen view lists all files... | 1. PDF output to be redesigned to list all documents in relevant ER sections 2. recode the exporter so it also exports the more generic documents attached to multiple Technical Files. 3 Explanatory documents and notes for this section will be added if applicable | 2017-11-01 | Derek Lamb
Oxygen Sensors | CE 97289 |  |  | Connection to other devices: It is unclear form the instructions for use / labeling how the end user determines proper connection to other devices. Precautions for change in performance: Incorrectly N/A by the manufacturer. The device is capable of performance issue which may require action or contact of the manufacturer. No instruction is provided.
Oxygen Sensors | CE 97289 |  |  | Manufacturer referenced documents 17847 and 17842 safety data sheets have not been provided. Additionally it is unclear if these sheets are provided to the end user. As such no evidence has been provided that performance is provided to the user.

## Sheet2
NonConID | RootCause | CorrectiveAction | CorrectiveActionTimeScale | CorrectiveActionResponsibility
| 1. Fault with PDF version of the Essential Requirement 2. File Exporter only Exporting Files directly tagged to Specific Technical files – The generic files that are used in multiple Technical files were not exporting. Reason not detected before. First remote audit. And First time PDF of Essential requirements used – digital on screen view lists all files... | 1. PDF output to be redesigned to list all documents in relevant ER sections 2. recode the exporter so it also exports the more generic documents attached to multiple Technical Files. 3 Explanatory documents and notes for this section will be added if applicable | 2017-11-01 | Derek Lamb
| Device has been on the market for 20+ Years there has not been a single reported case of ingress to the device. | None
| Firmware has not changed since the original design circa 1999. Zero reports of firmware failure in use. Firmware is NOT user upgradable | No further action required