QC 21 Non Conformance Report

| Date | 22.10.2020 |
| --- | --- |
| Issue id unique identifier | ### 201668 |
| BSI Ref (if applicable) unique identifier | 1970716-202010-N2 |
| Responsibility Person Overall responsible | Derek Lamb |
| Non-Conformance statement of the problem | The statement of the problem: The internal audit process isn't fully effective as no evidence were provided for the internal auditor to have been trained to the ISO 13485 version 2016. |
| Investigation By: Person responsible | Derek Lamb / Helen Lamb |
| Investigation Issue id (if applicable) Root Cause Analysis | 201674 In the training records there is no specific training for auditing of ISO 13483:2016. Audit training is present but not specific to 13485:2016. Root Cause Analysis When the auditor has done audit training in the past and the training record filled in, on the Training Manager, not enough detail has been added. So no details as to what version of the ISO standard training has been carried out. |
| Corrective Action By: Person responsible | Derek Lamb |
| Corrective Action Issue ID (if applicable): Relevant and Proportionate Corrective Action | #201708 |
| Time Scale for Corrective Action Time for completion of all identified actions | Immediate Action and Corrective Action to be carried out by 22nd November 2020 |
| Corrective Action: | Immediate Action To update the training manager with details of the last training carried out in house. Corrective Action To make sure more details are added to the training as it is entered in to the Training Manager. |
| Follow-up future issue id (Effectiveness verification) | #201708 |
| Effectiveness verification | To be able to look at the training carried out and be able to see specifics of what has been done. |
| Closed By: |  |
| Closed on |  |