QC 21 Non Conformance Report

| Date | 22.10.2020 |
| --- | --- |
| Issue id unique identifier | ### 201671 |
| BSI Ref (if applicable) unique identifier | 1970716-202010-N3 |
| Responsibility Person Overall responsible | Derek Lamb |
| Non-Conformance statement of the problem | The statement of the problem: The verification of purchased products is not fully effective as the sampling size of incoming goods wasn't seen to have been clearly defined. The organization explained that due to Covid19 pandemic they needed to handle a wide amount of incoming goods and a sampling method was used however no deviation was seen to have been documented in this regard |
| Investigation By: Person responsible | Derek Lamb / Helen Lamb |
| Investigation Issue id (if applicable) Root Cause Analysis | ### 201675 Our system has different levels of QA required on our incoming stock, depending on the supplier and type of stock. Due to the Covid 19 global pandemic we realised we had to get 50x more stock though the company, than we had before. Sales 2019 Qty sold 870 Units. Sales 2020 Qty Sold due to pandemic 49,694 Units. The need for the pandemic stock was immediate. Any delays to the unit from the supplier to the end user would result in less Ventilators in immediate use and almost certainly the loss of lives. Since we are not required to QA any of the stock we buy in to sell, as the stock is fully CE from the Supplier, The decision was made to Batch QA some of the stock coming in, but we did not update the procedures to reflect this. We also didn’t include a statement about this in the COVID 19 action plan. The Computer records which indicate type of QA - No QA,Full QA or Batch QA were updated by management to reflect the decision – which then allows product to be realsed Root Cause Analysis When we updated the system to allow batch QA of some products, when Covid 19 became an issue in March 2020, we did not update the relevant VOP or procedures. |
| Corrective Action By: Person responsible | Derek Lamb |
| Corrective Action Issue ID (if applicable): Relevant and Proportionate Corrective Action | #201710 #201713 |
| Time Scale for Corrective Action Time for completion of all identified actions | Immediate Action and Corrective Action to be carried out by 22nd November 2020 |
| Corrective Action: | Immediate Action To update the relevant VOP and Covid 19 action plan to reflect the decisions made about batch testing and sample size. Corrective Action We will update the Covid 19 action plan to cover the worse case scenarios for when emergency measures need to be taken. Make sure a Covid 19 or other extreme national or international crisis plan, is referred to in all tasks and mini audits. So that we have a record of our decisions and a plan of how to move forward, for each task, in an extreme national or international crisis. |
| Follow-up future issue id (Effectiveness verification) |  |
| Effectiveness verification |  |
| Closed By: |  |
| Closed on |  |